Alternative Energy Engineering

1155 Redway Drive, P.O. Box 339, Redway, CA 95560

(707)923-2277  FAX (707)923-3009

# Digital Amp Meter

28.9257 / 28.9259 / 28.9261

This meter may be used to monitor DC current on systems with voltages up to 65 VDC.  Accuracy is +/- 1% at 1/10 amp resolution with a 100 Amp shunt (P1 shorted) and 1 amp resolution when a 500 amp shunt is used (P1 removed).  Current draw of the meter is 0.008 Amps

Wiring Instructions:

Connect the B+ wire to positive voltage source.

Connect the B- wire to negative voltage source.

Connect the S1+ wire to shunt.

Connect the S1- wire to shunt (negative voltage source).

Wiring Notes:

The B- and the S1- are internally connected.  Care should be used to insure that the S1- of the meter is connected to the same side as B- of the meter on the shunt.

Shunt:

The shunt must be located on the negative cable of the power source between the source and the load.  A 100 amp 100mv shunt or a 500 amp 50mv shunt can be used if the jumper P1/P2 is installed correctly.

P1 & P2 Jumper:

The P1 and P2 jumpers are located the meter board.  For a 100 amp (1:1) shunt the jumper should be in the P1 location (factory default for 28.9257 & 28.9259).  For a 500 amp (10:1) shunt the jumper should be removed.  P2 is for future use.

Calibration:

This meter has been calibrated to be within 1.0 amp at 100 amp. It may be recalibrated by comparing it’s reading to that of another good quality amp meter.  Turning the white potentiometer on the circuit board behind the meter will allow you to adjust this meter to match your reference meter.

One Year Warranty:

Alternative Energy Engineering, Inc. warrants to the user that this unit is free from defects in materials and workmanship for the period of one year from the date of original purchase. If the unit fails within this period, it will be repaired or replaced at AEE's option, at no charge, when returned to the factory with a proof of purchase. This warranty does not apply to any product damaged by improper installation, accident, misuse, abuse, or excessive voltage, or if the product was changed or repaired by anyone other than AEE.  Alternative Energy Engineering, Inc. will not be liable for any consequential damages resulting from use or misuse of this product.  This warranty gives you specific legal rights, and you may have additional rights, which vary from state to state.

03/17/04 Rev 1.2